## Slide 1
Incorporate functional electrical stimulation with training system created by [3]
Test on healthy able-bodied participants to finalize the system
Rehabilitation Engineering Laboratory
Institute of Biomaterials and Biomedical Engineering
University of Toronto
Toronto Rehabilitation Institute | University Health Network
k.masani@utoronto.ca
www.toronto-fes.ca
DEVELOPMENT OF VISUAL FEEDBACK TRAINING FOR FUNCTIONAL ELECTRICAL STIMULATION THERAPY FOR BALANCE REHABILITATION
Kelvin Chow1, Emerson Grabke1, Jaeeun Yoo1, Jae Lee1, Kristin Musselman2, Kei Masani1,2
1) Institute of Biomaterials and Biomedical Engineering, University of Toronto
2) Toronto Rehabilitation Institute, University Health Network
Background
Training System Design
Develop a visual feedback training (VFT) system to integrate with functional electrical stimulation (FES) therapy
Individuals with spinal cord injuries often haven balance instability caused by reduced control of lower-limbed muscles
Rehabilitation tools have shown improvements in balance stability
  Visual Feedback Training [1]
  Functional Electrical Stimulation Therapy [2]
Hardware Design
CENTRE OF PRESSURE FEEDBACK
Force plate acquires 16 analog signals passed through NI DAQ; used to calculate centre of pressure (COP)
Visual COP feedback displayed on monitor (LabVIEW)
ELECTRICAL MUSCLE STIMULATOR
4 surface electrodes placed on participant’s shank muscles (tibialis anterior and gastrocnemius)
Electrodes control only anterior-posterior motion
Stimulation current controlled by PD controller using Compex muscle stimulator system
Completed by [3]
REHABILITATION GAMES
TRAINING PROTOCOL
Duration: 1 hour (30-40 minutes of standing time)
Static and dynamic balance assessments
Rehabilitation-based training games
STATIC BALANCE ASSESSMENT
DYNAMIC BALANCE ASSESSMENT
Assessment Performance Indicators
Root mean square distance – AP (RMS-AP)
Root mean square distance – ML (RMS-ML)
95% Confidence Ellipse Area (95% CE AREA)
Design Validation
CORRELATION BETWEEN GAME AND ASSESSMENT PERFORMANCE
REST PROTOCOL TO MITIGATE FATIGUE
10 healthy able-bodied participants completed one training session to gather data to validate the training system
2 participants completed 5 sessions each to test effectiveness of VFT over multiple training sessions
EFFECTIVENESS OF MULTIPLE TRAINING SESSIONS
A finalized visual feedback training system has been designed and validated, ready for clinical use
This system aims to help individuals with incomplete spinal cord injuries improve their balance ability, thus improving their quality of life
Next Steps
References
GUI Display
Dynamic Balance Area (DB AREA)
Purpose
Conclusion
TARGET
HUNTING
COLOUR MATCHING
ELLIPSE
Participants chose whether or not to take breaks in between games
Participants under voluntary rest were affected by fatigue
[1] Sayenko et al., Spinal Cord (2010) vol. 48, pp. 886-893
[2] Rouhani et al., Front Neurosci. (2017) 11:347
[3] Grabke et al., UnERD 2017
Participants show slight improvements over multiple training sessions in game and assessment performance
3 of 4 games show significant correlation to static balance ability, suggesting participants playing these games can improve static balance
VOLUNTARY REST
Participants forced to take a minimum of 6 breaks during session (>1 min.)
Participants under mandatory rest were not affected by fatigue
Current training system protocol implements mandatory rest periods
MANDATORY REST
GAME PERFORMANCE
DYNAMIC BALANCE ASSESSMENT
STATIC BALANCE
ASSESSMENT
TARGET
ELLIPSE
HUNTING
COLOUR
RMS-AP (mm)
RMS-ML (mm)
95% CE AREA (mm2)
DB AREA (mm2)
r=-0.87*
r=-0.79*
r=-0.84*
r=-0.89*
r=-0.84*
r=-0.92*
r=-0.01
r=-0.16
r=-0.08
r=0.18
r=-0.07
r=0.20
r=-0.58
r=-0.72
r=0.10
r=-0.78*
p=0.019*
p=0.27
p=0.12
p=0.21
p=0.54
p=0.73
Notes: Include hardware design
Swap methods and design